## Slide 1: Accurate Estimation of the volume flow rate in the cerebral artery using 3D cine phase-contrast MRI (4D-Flow)
Tokyo Institute of Technology |  | R. Ito, Y. Onishi, K. Amaya,
R'Tech Co. Ltd. |  | T. Kosugi, T. Kosugi,
Nagoya Univ. |  | H. Isoda, K. Ichikawa,
Hamamatsu Univ. Hosp. |  | Y. Takehara
Notes: Thank you
Today, I‘d like to talk to you about Accurate Estimation of the volume flow rate in the cerebral artery using 3D cine phase-contrast MRI

## Slide 2: Introduction
Target of our research
Patient-specific vascular computational fluid dynamics (CFD) for intracranial aneurysms
Goal of our research
Accurate calculation of WSS, pressure, etc. in aneurysms with CFD and making a tool for clinical quantitative hemodynamic evaluation
Notes: At the beginning, I'd like to show you the contents my presentation.
The Target of our research is Patient-specific vascular computational fluid dynamics (CFD) for intracranial aneurysms.
And Goal of our research is Accurate calculation of WSS, pressure, etc. in aneurysms with CFD, which is Tool for clinical quantitative hemodynamic evaluation.

## Slide 3: Introduction (cont.)
Three requirements for patient-specific CFD
  Accurate Vascular Shape,
  Accurate Blood Viscosity Model,
  Accurate Inlet and Outlet Boundary Conditions (BCs).
In this study, we focus on achieving “Accurate Inlet and Outlet BCs” using 4D Flow velocimetry.
Siemens scanner (Magnetom Verio 3.0T; Siemens AG, Healthcare Sector, Erlangen, Germany)
Notes: Vascular CFD analysis is believed to lose accuracy through 3 major routes.
1 vessel shape extraction
2 blood viscosity modeling
3 inlet and outlet boundary condition (BC) determinations.
In this study, we focus on 3. Accurate Determination of Inlet/Outlet BC.

## Slide 4: Problem in 4D Flow Velocimetry
For example…
Settings of a typical phantom study
The error of 4D Flow velocimetry is too large to ignore.
Target phantom
(acryl tube surrounded by agar)
Steady laminar flow of glycerol-water solution in the tube is measured.

## Slide 5: Problem in 4D Flow Velocimetry (cont.)
Velocity profile of a typical phantom study
Outside of tube (Agar)
Inside of tube
Acryl wall
Artifact 1:
Dragging velocity in the outside
Artifact 2: Lower velocity peak at the center
Cause of these artifacts: Phantom? 4D Flow itself? Or the both?
Anyway, it is difficult to obtain accurate velocity profile.
Acquisition Voxel Size: 1
These artifacts are NOT due to the partial volume effect.

## Slide 6: Strategy & Objective
Strategies for BC determination
Estimation of the volume flow rate (VFR) is a lot easier than that of the velocity profile.
Accurate VFR is a sufficient BC for CFD.
Objective
Propose a new method to estimate accurate volume flow rate (VFR)
using 4D-Flow velocimetry
Notes: Our purpose of this presentation is to Propose a new method to estimate accurate VFR using 4D-Flow velocimetry without accurate vessel shapes.

## Slide 7
Method:
Procedure of our
VFR estimation method

## Slide 8: Points
The average velocity error of 4D Flow measurement is nearly zero [Y.Onishi et al., IJNMBE, 2013].
The extravascular region is stationary.
The mean velocity in the extravascular region is nearly zero.
Area of extraction

## Slide 9: Procedure (1 of 3)
Measurement of the flow velocity vectors in all target vessel domains using 4D Flow.
Creation of 3D voxel data by combining the 4D Flow velocity magnitude images.

## Slide 10: Procedure (2 of 3)
Extract a vessel region to be larger than the exact vessel shape using the region growing method.
Conversion to a polygon data.
Extract this region
Exact vessel
Cross-section
view

## Slide 11: Procedure (3 of 3)
Configure many virtual cross-sections near the BC face.
Calculate VFR on each virtual cross-section .
Calculate the average of s, .
  We use as the VFR BC
BC face
VFR ()

## Slide 12
Result:
Validation experiments

## Slide 13: Experiment Device
Measurement of Hagen-Poiseuille flow in the tube
Straight tubes of =3.1, 6.5 mm
Water solution of glycerol of 40wt% (no contrast agents)
Steady laminar flow made by steady flow pump
Coriolis flowmeter measures the actual VFR
A thin-walled polyimide tube/agar phantom (thickness: 0.05 mm)

## Slide 14: Result 1: A Large Diameter Tube (=6.5 mm)
Actual VFR [mm3/s] | 10068
Estimated VFR [mm3/s] | 10190
Error | +1.9%
The VFR profile
The velocity profile
VFR is successfully estimated within 2% error.
Our new phantom could reduce both the dragging velocity and lower velocity peak artifacts.
The estimated VFR agreed with the theoretical curve.

## Slide 15: Result 2: A Small Diameter Tube (=3.1 mm)
Actual VFR [mm3/s] | 2496
Estimated VFR [mm3/s] | 3235
Error | +29.4%
The VFR profile
The velocity profile
VFR is estimated with 30% error, which is too large.
The two artifacts were reduced but there still remains non-negligible measurement errors.
The estimated VFR ultimately exceeded the exact VFR.
There remains a possibility of influence from the plastic wall.
Dispersion may occur in 4D Flow velocimetry itself.

## Slide 16
Summary

## Slide 17: Summary
A new method to estimate accurate VFR using 4D Flow velocimetry is proposed.
In case of steady laminar flow, the VFR in a tube > 6 mm in diameter can be accurately estimated by our method. But, in a tube < 6 mm in diameter, the VFR cannot be estimated with a practically sufficient accuracy.
It is NOT recommended to use thick-walled sold phantoms or bulk sold phantoms for 4D Flow validation tests.
Thank you for your attention.
And I appreciate your questions and comments in slow English without medical terms!!

## Slide 18
Appendix

## Slide 19: Acquisition Parameters (acryl)
Scanner | Siemens Magnetom Verio 3.0T
Coil | 12 ch Head coil
PAT | 2
Phase partial Fourier | 6/8
Slice partial Fourier | 6/8
TR [ms] | 33.6
TE [ms] | 4.32
Acquisition Time | 8:40
FOV [mm] | 160×160
Matrix | 160×160
Slice Thickness [mm] | 1.00
FA [deg] | 15
VENC [m/sec] | 1.2

## Slide 20: Acquisition Parameters (polyimide)
Scanner | Siemens Magnetom Verio 3.0T
Coil | 12 ch Head coil
PAT | Off
Phase partial Fourier | Off
Slice partial Fourier | Off
TR [ms] | 37.04
TE [ms] | 5.06
Acquisition Time | 26:41
FOV [mm] | 160×160
Matrix | 160×160
Slice Thickness [mm] | 1.00
FA [deg] | 15
VENC [m/sec] | 1.2

## Slide 21: Comparison of the maximum flow velocity
Case name | Exact maximum velocity [mm/s] | Measured maximum velocity [mm/s] | Error [%]
Acryl 3.0 mm | 600.12 | 346.68 | -42.2
Acryl 6.0 mm | 599.95 | 458.77 | -23.5
PL/Agar 3.1 mm | 661.38 | 511.34 | -22.7
PL/Agar 6.5 mm | 606.85 | 516.95 | -14.8
10-20% reduction in the lower velocity peak artifact.

## Slide 22: Validation Experiment on 1 Healthy Volunteer
Result with Siemens Scanner
Result with GE Scanner
We estimated average VFR in 1 heartbeat on 3 cross-sections (ICA, MCA and ACA).
The VFR becomes constant at the point where it exceeds the exact radius (approximately 2.5 mm).

## Slide 23: Points of Our New Phantom
Important factors for the phantom
The continuity of the proton density between the inner and outer areas of the tube.
The phantom moisture content is similar to the moisture content of the white matter of the brain.
Acquisition Voxel Size: 1
Polyimide tube wall thickness: 0.05 mm
It is difficult for the scanner to detect the thickness if the ratio to the spatial resolution is < 0.1.

## Slide 24: Extraction of Vessel Shape using 4D Flow
3D TOF MRA
A comparison of extracted silicone cerebral aneurysm phantom vessel shapes.
  Accurate extraction of vessel shapes using 4D Flow is difficult.
4D Flow rephased imaging
A number of surface deterioration point were included.
No vessel shapes were extracted (at arrow point).

## Slide 25: Review: Error Eval. of 4D-Flow
Histogram of error distribution of 4D Flow velocimetry
Each error of flow velocity components has similar distribution (1.73 mm/s).
Using a velocity profile in outside of the tube.
The average velocity error of measurement is almost zero.

## Slide 26: Result: Acryl Tube
 | 3.0 mm in inner diameter | 6.0 mm in inner diameter
Actual VFR [mm3/s] | 2121.0 | 8482.0
Estimated VFR [mm3/s] | 4318.0 | 10167
Error | +103.6% | +19.8%
The VFR profile
3.0 mm in inner diameter
6.0 mm in inner diameter
The VFR profile

## Slide 27: Procedure: Estimation of VFR
Estimation procedure
  Correct s of all inlets/outlets so that the sum of VFRs is exactly zero.
  e.g.)
  Set the corrected VFR () as the estimated VFR.
Constraint
Cost Function